CEN-Tech Multimeter

[image: image1.png]General

These 83 series instruments are a series of
compact pocket-sized 3 ' digital multimeters for
measuring DC and AC Voltage, DC Current,
Resistance and Diode. Some of those also
provide Temperature, Transistor measurement
and audible continuity test function or can be
used as a signal generator(see table). Full range
overload protection and low battery voltage
indication are provided. They are ideal
instruments for use in fields, such as laboratory,
workshop, DIYers and home applications.

M83 Series Multimeters

instrument.
To extend the life of this battery, the switch
should be in the "OFF" position when the
instrument is not in use.

2. DISPLAY
3 ' digit,7 segment, 0.5" high LCD.

3. "Common" JACK
Plug in connector for black (negative) test
lead.

4. "VQ mACx" JACK

Plug in connector for red (Positive) test lead
for all voltage and resistance and current
(except 10A) measurements.

5. ™O0A" JACK

Model| DCVIACV]DCA{OHMI-PHL D hPE (BAT LA L T RT) Plug in connector to red (positive) test lead
830A| v V v v v v
for 10A measurement.
83B| v | v | v | v |v v
v V' v v v |Iv] v v

L v v B o e B e - SPECIFICATIONS

81| v| v | v ]| v ]|v]~ Accuracies are guaranteed for 1 year, 23°C*t5 |
82 | v| v| v v v]iv]lv v less than 80%RH

833 V. v v v v | v v v

85| v|vIv]iv]v]y v v % DC VOLTAGE

87| v| v v] vy Vv RANGE [RESOLUTION ACCURACY.

88| v| v]v]v]v]vlv v 200mV 100UV +0.5% of rdg +2D
*RT: Room Temperature 2000mV 1mV +0.5% of rdg +3D

20V 10mV +0.5% of rdg %3D

FRONT PANEL DESCRIPTION 200V 100mV. +0.5% of rdg +3D
1. FUNCTION AND RANGE SWITCH 1000V, 1V +0.8% of rdg +3D

This switch is used to select the function and
desired range as well as to turn on the

OVERLOAD PROTECTION: 220V rms AC for
200mV range and 1000V DC or 750V rms for all




[image: image2.png]ranges. :

AC VOLTAGE
RANGE | RESOLUTION ACCURACY
200V 100mV.
+ +
750V v +1.2% of rdg +10D

RESPONSE: Average responding, calibrated in
rms of a sine wave.

FREQUENCY RANGE: 45Hz ~ 450Hz
OVERLOAD PROTECTION: 1000V DC or 750V
rms for all ranges.

DC CURRENT
RANGE | RESOLUTION ACCURACY
200uA 100nA + 1% of rdg £3D
2000uA 1uA +1% of rdg +3D
20mA 10uA =+ 1% of rdg £3D
200mA 100uA +1.5% of rdg * 5D
10A 10mA +2% of rdg +10D

OVERLOAD PROTECTION: 200mA 250V fuse
(10A range unfused).
MEASURING VOLTAGE DROP:200mV

RESISTANCE
RANGE | RESOLUTION ACCURACY
200Q 100mQ =+ 1.0% of rdg + 10D
2000Q 10 =+ 1.0% of rdg + 2D
20KQ 10Q =+ 1.0% of rdg + 2D
200KQ 100Q =+ 1.0% of rdg + 2D
2000KQ 1KQ =+ 1.0% of rdg * 2D

MAXIMUM OPEN CIRCUIT VOLTAGE: 3.2V
OVERLOAD PROTECTION: 15 seconds maxi-
mum 220Vrms.

e AUDIBLE CONTINUITY
RANGE DESCRIPTION
.))) Built-in buzzer sounds if resistance is|
less then 30+ 20Q
OVERLOAD PROTECTION: 15 second maxi-
mum 220 V rms.

TEMPERATURE (K TYPE PROBE)

RANGE | RESOLUTION| ACCURACY
+3°C*2D

20°C to 6 (up to 150°C)
1370 C +3% of rdg

(over 150°C)

OPERATING INSTRUCTIONS

WARNING

& Toavoid electrical shock hazard and/or
damage of the instrument, do not
measure voltages that might exceed
500V above earth ground.

&2 Before the use of instrument, inspect test
leads, connectors and probes for cracks,
breaks, or crazes in the insulation.




[image: image3.png]DC VOLTAGE MEASUREMENT

1. Connect red test lead to"VQ mA"jack, Black
lead to "COM"jack.

2. Set RANGE switch to desired DCV position,
if the voltage to be measured is not known
beforehand, set switch to the highest range
and reduce it until satisfactory reading is
obtained.

3. Connest test leads to device or circuit being
measured.

4. Turn on power of the device or circuit being
measured voltage value will appear on
Digital Display along with the voltage
polarity.

AC VOLTAGE MEASUREMENT

1. Red lead to "VQ mA". Black lead to "COM"
(for measurements between 200mA and
10A connect red lead to "10A"jack with fully
depressed.)

2. RANGE switch to desired DCA position.

3. Open the circuit to be measured, and
connect test leads INSERIES with the load
in with current is to measure.

4. Read current value on Digital Display.

RESISTANCE MEASUREMENT
1. Redlead to"VQ mA".Black lead to "COM".
2. RANGE switch to desired OHM position.

3. If the resistance being measured is connec-
ted to a circuit, turn off power and discharge
all capacitors before measurement.

4. Connect test leads to circuit being mea-
sured.

5. Read resistance value on Digital Display.

DIODE MEASUREMENT

1. Redlead to"VQ mA" Black lead to "COM".

2. RANGE switch to "PH" position.

3. Connect the red test lead to the anode of the
diode to be measured and black test lead to
cathode.

4. The forward voltage drop in mV will be
displayed. If the diode is reversed, figure "1"
will be shown.

TRANSISTOR hFE MEASUREMENT

1. RANGE switch to the hFE position.

2. Determine whether the transistor is PNP of
NPN type and locate the Emitter, Base and
Collector leads. Insert the leads into the
proper holes of the hFE Socket on the front
panel.

3.  The meter will display the approximate hFE
value at the condition of base current 10uA
and V¢e2.8V.

TEMPERATURE MEASUREMENT
1. Connect the K type thermoelectric couple to

"VQ mA" and "COM" jacks.
2. RANGE switch to TEMP position.
3. The display will read Temperature value ‘C

AUDIBLE CONTINUITY TEST

1. Redleadto"VQ mA" Black lead to "COM".

2. RANGE switch to"*) "position.

3. Connect test leads to two points of circuit to
be tested. If the resistance is lower then 30Q

+20Q, the buzzer will sound.

TEST SIGNAL USE
1. RANGE switch to "JU™" position.

2. Atest signal (50Hz) appears between "VQ
mA" and "COM" jack, the output voltage is
approx 5V p-p with 50KQ impedance.

BATTERY AND FUSE REPLACEMENT

Fuse rarely need replacement and blow almost

always as a result of operator error.

If "E2" appears in display, it indicates that the
battery should be replaced.

To replace battery & Fuse (200mA/250V)remove
the 2 screws in the bottom of the case, simply

remove the old. and replace with a new one. Be
careful to observe polarity.




